## Method
Shielding Design for Diagnostical X-ray Facilities
Authors: | Pedro Ruiz Manzano and Maria Angeles Rivas Ballarín
Hospital Clínico Universitario "Lozano Blesa"
Avda San Juan Bosco 15. 50009-Zaragoza
pruizm@salud.aragon.es
mrivasb@salud.aragon.es
Calculation Method
Primary Beam |  | A = (G. W.T.U)/(d2.Hw)
Scattered Radiation |  | A = (G. W.T.U.a.S)/(dp2.ds2.400.Hw)
Leakeage Radiation: |  | A = (F. W.T.U)/(d2.Qh.Hw)
Where:
A: | Attenuation Factor
W: | Workload in mA.min/week
Qh : | Maximum tube load, in mA.min/h ( 300 for 100 kVp, 240 for 125 kVp, 200 for 150 kVp)
T: | Occupancy Factor
U: | Use factor
d: | Focus-barrier distance
dp: | Focus-patient distance
ds: | Patient-secondary barrier distance
Hw: | Weekly Dose Limit for the area
a: | Scattering Factor (0,002 for a 400 cm2 field on patient)
S: | Field size
G: | Tube Output
F: | Leakeage Radiation Correction Factor (most unfavorable value = 1)
Based on l Guia CSN 5.11 "Aspectos técnicos de seguridad y protección radiológica de instalaciones médicas de rayos X para diagnóstico" .This document can be downloaded at the CSN website:
http://www.csn.es/images/stories/publicaciones/guias_seguridad/5.11.pdf

## Primary barrier example
Example of Primary Barrier Calculation
Chest Bucky Wall A on a chest dedicated x-ray unit.
Facility: |  |  | Barrier: | Wall
Room: | 1 (Chest) |  | Adjacent room: | Sala espera
Date: |  |  | Sign on the room plan : | A
GENERAL DATA |  | Symbol | Values
Use factor |  | U | 1
Occupancy factor |  | T | 0.25
Area |  |  | Uncontrolled
Weekly dose limit (mSv) |  | Hw | 0.02
WORKLOAD DATA
Weekly workload (mA.min) |  | W | 80
Maximum operating potential (50-150) |  | KVp | 125
Tube output (mSv. m2/mA.min) |  | G | 11.1
DISTANCES
Focus-barrier distance (m) |  | d | 2
ATTENUATION FACTOR |  | A | 2775
LEAD THICKNESS (DIN 6812) (mm) |  |  | 2.005
Equivalent thickness of other materials
High density concrete (3,2 g/cm3) DIN 6812 (cm) |  |  | 2.5063
Concrete (2,3 g/cm3 ) DIN 6812 (cm) |  |  | 16.03
Clay brick (1,8 g/cm3) DIN 6812 (cm) |  |  | 22.7862
Note on C16: Fill only cells in blue
Note on F24: Beam always directed to this wall
Note on F30: Value given in Guía 5.11 l CSN for this kind of room
Note on F31: Typical value for chest exams.
Note on F37: A = (G. W.T.U)/(d2.Hw)
Note on F39: Interpolation in Fig 2 Guía 5.11 CSN.
Note on F42: Interpolation in tables DIN 6812
Waiting Room
2 m
A
B
Office Secretaría
Patient

## Secondary barrier example
Example of shielding calculation for secondary barrier
Wall B on a chest dedicated-ray unit, perpendicular to a primary barrier..
Facility: |  |  | Barrier: |  | Wall
Room: | 1 (Chest) |  | Adjacent room: |  | Office
Date: |  |  | Sign on the room plan : |  | B
GENERAL DATA |  | Symbol | Scatter values | Leakage values
Use factor |  | U | 1
Occupancy factor |  | T | 1
Area |  |  | Uncontrolled
Weekly dose limit (mSv) |  | Hw | 0.02
WORKLOAD DATA
Leakage maximum load (mA.min) |  | Qh |  | 240
Weekly workload (mA.min) |  | W | 80
Maximum operating potential (50-150) |  | KVp | 125
Tube output (mSv. m2/mA.min) |  | G | 11.1
Leakage correction factor |  | F |  | 0.68
Leakage half value layer (mm) |  | HVL |  | 0.27
DISTANCES
Focus-patient distance (m) |  | dp | 1.6
Patient-barrier distance (m) |  | ds | 1.7
Focus-barrier distance (m) |  | d |  | 1.7
SCATTER PARAMETERS
Field size (cm2) |  | S | 1500
Scatter factor |  | a | 0.002
ATTENUATION FACTOR |  | A | 45.0097 | 3.9216
LEAD THICKNESS (DIN 6812) (mm) |  |  | 0.65508 | 0.53229
TOTAL LEAD THICKNESS (mm Pb): |  |  | 0.92508
Equivalent thickness of other materials
High density concrete (3,2 g/cm3) DIN 6812 (cm) |  |  | 1.0878
Concrete (2,3 g/cm3 ) DIN 6812 (cm) |  |  | 8.6431
Clay brick (1,8 g/cm3) DIN 6812 (cm) |  |  | 12.6309
Note on C17: Fill only cells in blue
Note on F23: Use factor secondary barrier
Note on F44: Proposed value in Guía 5.11 CSN for S= 400 cm2
Note on F46: A = (G.W.T.U.a.S)/(dp2.ds2.400.Hw)
Note on F48: Interpolation in Fig 2 Guía 5.11 CSN.
Note on F50: If thicknesses for scattered and leakage radiation differ more than 1 leakage TVL, the bigger thickness is taken. If they differ less than 1 leakage TVL, a leakage HVL ia added to the bigger thickness.
Note on F52: Interpolation in tables DIN 6812
Note on G29: Proposed value in Guía 5.11 CSN
Note on G33: Value according to Fig 3 Guía 5.11 CSN
Note on G34: Value according to Table 4 Guía 5.11 CSN
Note on G46: A = (F. W.T.U)/(d2.Qh.Hw)
Note on G48: E= CHR·lnA/ln2
Office
Waiting room
2 m
A
B
Patient
1,7 m
1,6 m
1,7 m

## Primary barrier calculation
Rendimiento (mSv.m2/mA.min)
PRIMARY BARRIER CALCULATION |  |  | Indice |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 |  | KV1 | KV2
|  |  | 1 | Kvp | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 |  | 8 | 9
Facility: |  | Barrier: | 2 | G | 1.8 | 2.8 | 3.8 | 5 | 6.2 | 7.5 | 9 | 10.2 | 12 | 13 | 14 |  | 120 | 130 | [11 empty] | Pared
Room: |  | Adjacent room: | [14 empty] | 10.2 | 12 | [11 empty] | Pasillo
Date: |  | Sign on the room plan : | [27 empty] | A
Blindaje (mm Pb)
|  |  | Indice | 1 | 1 50: 2; 2 60: 3; 3 70: 4; 4 80: 5; 8 120: KV1; 9 130: KV2
GENERAL DATA | Symbol | Values | 1 | A/KVp | 1 50: 50; 2 60: 75; 3 70: 100; 4 80: 150; 8 120: 100; 9 130: 150 |  |  |  | Hormigón baritado
Use factor | U | 1 | 2 | 1 | 0 | 0 | 0 | 0 |  |  | Indice | 4 | 5 |  | Int.1 | Int |  | indice | 1 | 2 | 3 | 4 |  |  |  | KV1 | KV2
Occupancy factor | T | 0.25 | 3 | 3 | 0.03 | 0.06 | 0.1 | 0.17 | A1 | 1000 | 14 | 1.5 | 1.8 |  | 1.65 |  |  | 1 | mmPb/KV | 50 | 100 | 150 |  |  |  | 100 | 150
Area |  | Libre acceso | 4 | 5 | 0.04 | 0.09 | 0.16 | 0.26 | A2 | 3000 | 15 | 1.8 | 2.3 |  | 2.05 | 2.005 |  | 2 | 0.5 | 1.5 | 0.4 | 0.73 |  |  | Indice | 3 | 4 | Int.1 | Int
Weekly dose limit (mSv) | Hw | 0.02 | 5 | 8 | 1 50: 0.05; 2 60: 0.12; 3 70: 0.22; 4 80: 0.34 |  |  | 3 | 1 | 3.1 | 0.86 | 1.5 | Pb1 | 2 | 4 | 1.7 | 3.3 | 2.5
|  |  | 6 | 10 | 1 50: 0.09; 2 60: 0.14; 3 70: 0.26; 4 80: 0.38 |  |  | 4 | 2 | 6.1 | 1.7 | 3.3 | Pb2 | 3 | 5 | 2.4 | 5.1 | 3.75 | 2.5063
WORKLOAD DATA |  |  | 7 | 30 | 1 50: 0.1; 2 60: 0.25; 3 70: 0.45; 4 80: 0.62 |  |  | 5 | 3 | 8.6 | 2.4 | 5.1
Weekly workload (mA.min) | W | 80 | 8 | 50 | 1 50: 0.13; 2 60: 0.31; 3 70: 0.59; 4 80: 0.8 |  |  | 6 | 4 | 10.7 | 3 | 6.7
Maximum operating potential (50-150) | KVp | 125 | 9 | 80 | 1 50: 0.15; 2 60: 0.38; 3 70: 0.71; 4 80: 0.9
Tube output (mSv. m2/mA.min) | G | 11.1 | 10 | 100 | 1 50: 0.16; 2 60: 0.41; 3 70: 0.8; 4 80: 1 |  |  |  | En negrita valores extrapolados
|  |  | 11 | 300 | 1 50: 0.25; 2 60: 0.6; 3 70: 1.1; 4 80: 1.4
DISTANCES |  |  | 12 | 500 | 1 50: 0.3; 2 60: 0.7; 3 70: 1.25; 4 80: 1.6 |  |  |  | Hormigón
Focus-barrier distance (m) | d | 2 | 13 | 800 | 1 50: 0.35; 2 60: 0.75; 3 70: 1.45; 4 80: 1.75 |  |  | indice | 1 | 2 | 3 | 4 |  |  |  | KV1 | KV2
|  |  | 14 | 1000 | 1 50: 0.4; 2 60: 0.8; 3 70: 1.5; 4 80: 1.8 |  |  | 1 | mmPb/KV | 50 | 100 | 150 |  |  |  | 100 | 150
ATTENUATION FACTOR | A | 2775 | 15 | 3000 | 1 50: 0.45; 2 60: 1.1; 3 70: 1.8; 4 80: 2.3 |  |  | 2 | 0.5 | 6.2 | 4.4 | 6 |  |  | Indice | 3 | 4 | Int.1 | Int
|  |  | 16 | 5000 | 1 50: 0.48; 2 60: 1.2; 3 70: 2.2; 4 80: 2.5 |  |  | 3 | 1 | 13 | 8 | 10.5 | Pb1 | 2 | 4 | 14 | 18 | 16
LEAD THICKNESS (DIN 6812) (mm) |  | 2.005 | 17 | 8000 | 1 50: 0.49; 2 60: 1.3; 3 70: 2.3; 4 80: 2.7 |  |  | 4 | 2 | 23 | 14 | 18 | Pb2 | 3 | 5 | 19 | 25 | 22 | 16.03
Equivalent thickness of other materials |  |  | 18 | 10000 | 1 50: 0.5; 2 60: 1.35; 3 70: 2.4; 4 80: 2.8 |  |  | 5 | 3 | 31 | 19 | 25
High density concrete (3,2 g/cm3) DIN 6812 (cm) |  | 2.5063 | 19 | 30000 | 1 50: 0.6; 2 60: 1.5; 3 70: 2.75; 4 80: 3.25 |  |  | 6 | 4 | 39 | 24 | 30
Concrete (2,3 g/cm3 ) DIN 6812 (cm) |  | 16.03 | 20 | 50000 | 1 50: 0.65; 2 60: 1.6; 3 70: 3; 4 80: 3.5
Clay brick (1,8 g/cm3) DIN 6812 (cm) |  | 22.7862 | 21 | 80000 | 1 50: 0.69; 2 60: 1.65; 3 70: 3.1; 4 80: 3.6 |  |  |  | En negrita valores extrapolados
|  |  | 22 | 100000 | 1 50: 0.7; 2 60: 1.7; 3 70: 3.2; 4 80: 3.7 |  |  |  | Ladrillo macizo
|  |  | 23 | 300000 | 1 50: 0.8; 2 60: 2; 3 70: 3.55; 4 80: 4.2 |  |  | indice | 1 | 2 | 3 | 4 |  |  |  | KV1 | KV2
|  |  | 24 | 500000 | 1 50: 0.85; 2 60: 2.15; 3 70: 3.7; 4 80: 4.4 |  |  | 1 | mmPb/KV | 50 | 100 | 150 |  |  |  | 100 | 150
|  |  | 25 | 800000 | 1 50: 0.89; 2 60: 2.25; 3 70: 3.9; 4 80: 4.55 |  |  | 2 | 0.5 | 10 | 7 | 8.4 |  |  | Indice | 3 | 4 | Int.1 | Int
|  |  | 26 | 1000000 | 1 50: 0.9; 2 60: 2.3; 3 70: 4; 4 80: 4.7 |  |  | 3 | 1 | 20 | 12 | 15 | Pb1 | 2 | 4 | 19.5 | 26 | 22.75
|  |  | 27 | 3000000 | 1 50: 0.95; 2 60: 2.5; 3 70: 4.4; 4 80: 5.2 |  |  | 4 | 2 | 33 | 19.5 | 26 | Pb2 | 3 | 5 | 26 | 34 | 30 | 22.7862
|  |  | 28 | 5000000 | 1 50: 1; 2 60: 2.65; 3 70: 4.6; 4 80: 5.4 |  |  | 5 | 3 | 44 | 26 | 34
|  |  | 29 | 8000000 | 1 50: 1.05; 2 60: 2.7; 3 70: 4.75; 4 80: 5.5 |  |  | 6 | 4 | 60 | 33 | 42
|  |  | 30 | 10000000 | 1 50: 1.1; 2 60: 2.8; 3 70: 4.8; 4 80: 5.6
Note on C4: Fill only cells in blue
Note on H4: Rellenar únicamente las celdas azules

## Secondary barrier calculation
Rendimiento (mSv.m2/mA.min)
SECONDARY BARRIER CALCULATION |  |  |  | Indice |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 |  | KV1 | KV2
|  |  |  | 1 | Kvp | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 |  | 8 | 9
Facility: |  | Barrier: |  | 2 | G | 1.8 | 2.8 | 3.8 | 5 | 6.2 | 7.5 | 9 | 10.2 | 12 | 13 | 14 |  | 120 | 130
Room: |  | Adjacent room: | [15 empty] | 10.2 | 12
Date: |  | Sign on the room plan :
Blindaje (mm Pb)
|  |  |  | Indice | 1 | 1 50: 2; 2 60: 3; 3 70: 4; 4 80: 5; 8 120: KV1; 9 130: KV2
GENERAL DATA | Symbol | Scatter values | Leakage values | 1 | A/KVp | 1 50: 50; 2 60: 75; 3 70: 100; 4 80: 150; 8 120: 100; 9 130: 150 |  |  |  | Hormigón baritado
Use factor | U | 1 |  | 2 | 1 | 0 | 0 | 0 | 0 |  |  | Indice | 4 | 5 |  | Int.1 | Int |  | indice | 1 |  | 2 | 3 | 4 |  |  |  | KV1 | KV2
Occupancy factor | T | 1 |  | 3 | 3 | 0.03 | 0.06 | 0.1 | 0.17 | A1 | 30 | 7 | 0.45 | 0.62 |  | 0.535 |  |  | 1 | mmPb/KV |  | 50 | 100 | 150 |  |  |  | 100 | 150
Area |  | Uncontrolled |  | 4 | 5 | 0.04 | 0.09 | 0.16 | 0.26 | A2 | 50 | 8 | 0.59 | 0.8 |  | 0.695 | 0.65508 |  | 2 | 0.5 |  | 1.5 | 0.4 | 0.73 |  |  | Indice | 3 | 4 | Int.1 | Int
Weekly dose limit (mSv) | Hw | 0.02 |  | 5 | 8 | 1 50: 0.05; 2 60: 0.12; 3 70: 0.22; 4 80: 0.34 |  |  | 3 | 1 |  | 3.1 | 0.86 | 1.5 | Pb1 | 0.5 | 2 | 0.4 | 0.73 | 0.565
|  |  |  | 6 | 10 | 1 50: 0.09; 2 60: 0.14; 3 70: 0.26; 4 80: 0.38 |  |  | 4 | 2 |  | 6.1 | 1.7 | 3.3 | Pb2 | 1 | 3 | 0.86 | 1.5 | 1.18 | 1.0878
WORKLOAD DATA |  |  |  | 7 | 30 | 1 50: 0.1; 2 60: 0.25; 3 70: 0.45; 4 80: 0.62 |  |  | 5 | 3 |  | 8.6 | 2.4 | 5.1
Leakage maximum load (mA.min) | Qh |  | 240 | 8 | 50 | 1 50: 0.13; 2 60: 0.31; 3 70: 0.59; 4 80: 0.8 |  |  | 6 | 4 |  | 10.7 | 3 | 6.7
Weekly workload (mA.min) | W | 80 |  | 9 | 80 | 1 50: 0.15; 2 60: 0.38; 3 70: 0.71; 4 80: 0.9
Maximum operating potential (50-150) | KVp | 125 |  | 10 | 100 | 1 50: 0.16; 2 60: 0.41; 3 70: 0.8; 4 80: 1 |  |  |  | En negrita valores extrapolados
Tube output (mSv. m2/mA.min) | G | 11.1 |  | 11 | 300 | 1 50: 0.25; 2 60: 0.6; 3 70: 1.1; 4 80: 1.4
Leakage correction factor | F |  | 0.68 | 12 | 500 | 1 50: 0.3; 2 60: 0.7; 3 70: 1.25; 4 80: 1.6 |  |  |  | Hormigón
Leakage half value layer (mm) | HVL |  | 0.27 | 13 | 800 | 1 50: 0.35; 2 60: 0.75; 3 70: 1.45; 4 80: 1.75 |  |  | indice | 1 |  | 2 | 3 | 4 |  |  |  | KV1 | KV2
|  |  |  | 14 | 1000 | 1 50: 0.4; 2 60: 0.8; 3 70: 1.5; 4 80: 1.8 |  |  | 1 | mmPb/KV |  | 50 | 100 | 150 |  |  |  | 100 | 150
DISTANCES |  |  |  | 15 | 3000 | 1 50: 0.45; 2 60: 1.1; 3 70: 1.8; 4 80: 2.3 |  |  | 2 | 0.5 |  | 6.2 | 4.4 | 6 |  |  | Indice | 3 | 4 | Int.1 | Int
Focus-patient distance (m) | dp | 1.6 |  | 16 | 5000 | 1 50: 0.48; 2 60: 1.2; 3 70: 2.2; 4 80: 2.5 |  |  | 3 | 1 |  | 13 | 8 | 10.5 | Pb1 | 0.5 | 2 | 4.4 | 6 | 5.2
Patient-barrier distance (m) | ds | 1.7 |  | 17 | 8000 | 1 50: 0.49; 2 60: 1.3; 3 70: 2.3; 4 80: 2.7 |  |  | 4 | 2 |  | 23 | 14 | 18 | Pb2 | 1 | 3 | 8 | 10.5 | 9.25 | 8.6431
Focus-barrier distance (m) | d |  | 1.7 | 18 | 10000 | 1 50: 0.5; 2 60: 1.35; 3 70: 2.4; 4 80: 2.8 |  |  | 5 | 3 |  | 31 | 19 | 25
|  |  |  | 19 | 30000 | 1 50: 0.6; 2 60: 1.5; 3 70: 2.75; 4 80: 3.25 |  |  | 6 | 4 |  | 39 | 24 | 30
SCATTER PARAMETERS |  |  |  | 20 | 50000 | 1 50: 0.65; 2 60: 1.6; 3 70: 3; 4 80: 3.5
Field size (cm2) | S | 1500 |  | 21 | 80000 | 1 50: 0.69; 2 60: 1.65; 3 70: 3.1; 4 80: 3.6 |  |  |  | En negrita valores extrapolados
Scatter factor | a | 0.002 |  | 22 | 100000 | 1 50: 0.7; 2 60: 1.7; 3 70: 3.2; 4 80: 3.7 |  |  |  | Ladrillo macizo
|  |  |  | 23 | 300000 | 1 50: 0.8; 2 60: 2; 3 70: 3.55; 4 80: 4.2 |  |  | indice | 1 |  | 2 | 3 | 4 |  |  |  | KV1 | KV2
ATTENUATION FACTOR | A | 45.0097 | 3.9216 | 24 | 500000 | 1 50: 0.85; 2 60: 2.15; 3 70: 3.7; 4 80: 4.4 |  |  | 1 | mmPb/KV |  | 50 | 100 | 150 |  |  |  | 100 | 150
|  |  |  | 25 | 800000 | 1 50: 0.89; 2 60: 2.25; 3 70: 3.9; 4 80: 4.55 |  |  | 2 | 0.5 |  | 10 | 7 | 8.4 |  |  | Indice | 3 | 4 | Int.1 | Int
LEAD THICKNESS (DIN 6812) (mm) |  | 0.65508 | 0.53229 | 26 | 1000000 | 1 50: 0.9; 2 60: 2.3; 3 70: 4; 4 80: 4.7 |  |  | 3 | 1 |  | 20 | 12 | 15 | Pb1 | 0.5 | 2 | 7 | 8.4 | 7.7
|  |  |  | 27 | 3000000 | 1 50: 0.95; 2 60: 2.5; 3 70: 4.4; 4 80: 5.2 |  |  | 4 | 2 |  | 33 | 19.5 | 26 | Pb2 | 1 | 3 | 12 | 15 | 13.5 | 12.6309
TOTAL LEAD THICKNESS (mm Pb): |  | 0.92508 |  | 28 | 5000000 | 1 50: 1; 2 60: 2.65; 3 70: 4.6; 4 80: 5.4 |  |  | 5 | 3 |  | 44 | 26 | 34
Equivalent thickness of other materials |  |  |  | 29 | 8000000 | 1 50: 1.05; 2 60: 2.7; 3 70: 4.75; 4 80: 5.5 |  |  | 6 | 4 |  | 60 | 33 | 42
High density concrete (3,2 g/cm3) DIN 6812 (cm) |  | 1.0878 |  | 30 | 10000000 | 1 50: 1.1; 2 60: 2.8; 3 70: 4.8; 4 80: 5.6
Concrete (2,3 g/cm3 ) DIN 6812 (cm) |  | 8.6431
Clay brick (1,8 g/cm3) DIN 6812 (cm) |  | 12.6309
CDR fuga
[5 empty] | Indice | 2 60: 1; 3 70: 2; 4 80: 3; 5 90: 4; 6 100: 5 |  |  | KV1 | KV2 | CDR f
[5 empty] | 1 | 1 50: Kvp; 2 60: 50; 3 70: 75; 4 80: 100; 5 90: 125; 6 100: 150 |  |  | 4 | 5
[5 empty] | 2 | 1 50: CDR; 2 60: 0.18; 3 70: 0.5; 4 80: 0.84; 5 90: 0.9; 6 100: 0.96 |  |  | 125 | 150
[19 empty] | 0.9 | 0.96 | 0.9
CDR dispersa
[5 empty] | Indice | 2 60: 1; 3 70: 2; 4 80: 3; 5 90: 4 |  |  | KV1 | KV2 | CDR f
[5 empty] | 1 | 1 50: Kvp; 2 60: 50; 3 70: 75; 4 80: 100; 5 90: 150 |  |  | 3 | 4
[5 empty] | 2 | 1 50: CDR; 2 60: 0.1; 3 70: 0.55; 4 80: 0.9; 5 90: 1 |  |  | 100 | 150 |  | siD>F | 0.92508
[19 empty] | 0.9 | 1 | 0.95 | siF>D
Note on C4: Fill only cells in blue
Note on G20: In tube housings designed for nominal 150 KVp : 1 for 150 KVp 0,85 for 140 KVp 0,68 for 130 KVp 0,5 for 120 KVp 0,3 for 110 KVp 0,18 for 100 KVp