## Method
Spreadsheet for patient dose estimation in mammography
María Angeles Rivas Ballarín and Pedro Ruiz Manzano.
mrivasb@salud.aragon.es
pruizm@salud.aragon.es
Robson's parametric method to determine output and HVL of mammographic tubes 1 .
By using a computer model, the author produces data that allow the calculation of tube output and half value layer for the range of clinically encountered conditions . To do that, the following equations are used.
Equation 1: | log 10 (ESAK) = nlog10(kV)+log10(A)
Equation 2: | HVL = a(kV)2+b(kV)+c
Where calculated values for a, b, n are:
Target/Filter |  | Filter thickness |  | n | a | b
Mo/30 mm Mo |  | 36,1 mm |  | 3.06 | -0.000326 | 0.0273
Mo/25 mm Rh |  | 29,9 mm |  | 3.24 | -0.000624 | 0.0445
Rh/25 mm Rh |  | 29,9 mm |  | 3.03 | -0.000514 | 0.0425
W/50 mm Rh |  | 58,9 mm |  | 1.96 | -0.000539 | 0.0403
Rh/1,0 mm Al |  | 1,20 mm |  | 4.39 | -0.00113 | 0.0909
Mo/1,0 mm Al |  | 1,20 mm |  | 4.23 | -0.000775 | 0.0593
The use of these parameters to determine the tube output and HVL is made as follows:
1- Measure the tube potential for a nominal value of 28 kV.
2- Measure the tube output and HVL at a nominal tube voltage of 28 kV for each target-filter combination, behind the compression plate.
3- Using the values found in steps 1 and 2, and the appropriate parameters from the Table, calculate the values of A and c from Equations 1 and 2.
4- With all the parameters now known, the tube output and HVL can be calculated for any other measured value of tube voltage in the range 25–32 kV.
Estimation of mean glandular breast dose using Dance's additional factors2
According to the method proposed by Dance in 1990, mean glandular breast dose was calculated using the equation:
MGD = K·g
where K was the Entrance Surface Air Kerma, and g a conversion factor depending on HVL, assuming a 50% glandular/adipose tissue composition in breast.
A new equation is proposed now: |  |  |  | MGD = K·g·c·s
where K and g are the same as before, the factor c corrects for any difference in breast composition from 50% glandularity and the factor s corrects for any difference from the original tabulation due to the use of a different x-ray spectrum.
Glandularity c-factors are tabulated for two age groups: 40-49 years and 50-64 years. Therefore, this equation cannot be used out of that age range.
s-factors are tabulated for the following spectra: Mo-Mo, Mo-Rh, Rh-Rh, Rh-Al, W-Rh.
Dose to typical breasts simulated with PMMA 4
The doses to a range of typical breasts can be assessed using blocks of PMMA as breast substitutes. It should be noted that since PMMA is generally denser than breast tissue any automatic selection of kV, target or filter may be slightly different from real breasts. On systems that determine the exposure factors on breast thickness, this can be corrected by adding expanded polystyrene blocks to the PMMA as a spacer to make up a total thickness equal to the equivalent breast.
The average glandular dose (D) to a typical breast of thickness and composition equivalent to the thickness of PMMA tested is calculated by applying the following formula.
MGD = K·g·c·s
where K is the entrance surface air kerma (without backscatter) calculated at the upper surface of the PMMA. The factor g, corresponds to a glandularity of 50%,and the c-factor is given for typical breasts in the age range 50 to 64 .
How to use the spreadsheet
a) Patient dose: For Mo-Mo, Mo-Rh, Rh-Rh target/filter combinations, select CalMammo_1. For the other combinations, choose CalMammo_2. Introduce patient data in the relevant sheet.
b) PMMA dose: For Mo-Mo, Mo-Rh, Rh-Rh target/filter combinations, select PMMA dosimetry_1. For the other combinations, choose PMMA dosimetry_2
Data corresponding to the mammography unit should also be included, as:
a) KVp, tube output and HVL at nominal 28 kVp for each target-filter combination.
b) Table containing nominal set KVp versus measured KVp .
- Entrance Surface Dose (ESD) is calculated multiplying ESAK by the retrodispersion factor Fr, HVL-dependent, which is tabulated in European protocol on Dosimetry in Mammography3.
References
1.Robson K J. "A parametric method for determining mammographic X-ray tube output and half value layer". Br J Radiol 74 (2001), 335-340.
2. Dance D R, Skinner C L, et al. " Additional factors for the estimation of mean glandular breast dose using the UK mammography dosimetry protocol". Phys. Med.Biol. 45 (2000) 3225-3240.
3. European protocol on Dosimetry in Mammography. EUR 16263 EN
4.The European Protocol for the Quality Control of the physical and technical aspects of mammography screening (Addendum on digital Mammography).Appendix 1

## CalMammo_1
Factor s
Patient dose estimation in mammography | [23 empty] | Mo-Mo | 1
[24 empty] | Mo-Rh | 1.017
| Craniocaudal Projection | [22 empty] | Rh-Rh | 1.061
Date | [23 empty] | Factor c grupo edad 40 a 49 años | [8 empty] | Factor c grupo edad 50 a 64 años | [20 empty] | Valor de g
[24 empty] | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | [15 empty] | 1 | 2 | 3 | 4 | 5 | 6
Facility: | [10 empty] | kVp set | kVpmeas | [11 empty] | Espesor | HVL (mm Al) | [7 empty] | Espesor | HVL (mm Al) | [7 empty] | Combinación ánodo/filtro |  | Espesor de filtro |  | n | a | b | A | c |  |  |  | Espesor | HVL (mm Al)
Room: | [10 empty] | 23 | 23 | [10 empty] | 1 | mama | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 |  |  | 1 | mama | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 |  |  |  | Mo/30 mm Mo |  | 36,1 mm |  | 3.06 | -0.000326 | 0.0273 | 0.0010921 | -0.14972 | Mo-Mo |  | 1 | mama | 0.3 | 0.35 | 0.4 | 0.45 | 0.5
Unit: | [10 empty] | 24 | 24 | [10 empty] | 2 | 2 | 0.885 | 0.891 | 0.9 | 0.905 | 0.91 |  |  | 2 | 2 | 0.885 | 0.891 | 0.9 | 0.905 | 0.91 |  |  |  | Mo/25 mm Rh |  | 29,9 mm |  | 3.24 | -0.000624 | 0.0445 | 0.00049788 | -0.31773 | Mo-Rh |  | 2 | 2 | 0.39 | 0.433 | 0.473 | 0.509 | 0.543
[11 empty] | 25 | 24.9 | [10 empty] | 3 | 3 | 0.894 | 0.898 | 0.903 | 0.906 | 0.911 |  |  | 3 | 3 | 0.925 | 0.929 | 0.931 | 0.933 | 0.937 |  |  |  | Rh/25 mm Rh |  | 29,9 mm |  | 3.03 | -0.000514 | 0.0425 | 0.00095827 | -0.33839 | Rh-Rh |  | 3 | 3 | 0.274 | 0.309 | 0.342 | 0.374 | 0.406
|  |  |  | Mo-Mo | Mo-Rh | Rh-Rh |  |  |  |  | 26 | 25.9 | [10 empty] | 4 | 4 | 0.94 | 0.943 | 0.945 | 0.947 | 0.948 |  |  | 4 | 4 | 1 | 1 | 1 | 1 | 1 | [14 empty] | 4 | 4 | 0.207 | 0.235 | 0.261 | 0.289 | 0.318
kV ( nominal) |  |  |  | 28 | 28 | 28 |  |  |  |  | 27 | 27 | [10 empty] | 5 | 5 | 1.005 | 1.005 | 1.005 | 1.004 | 1.004 |  |  | 5 | 5 | 1.086 | 1.082 | 1.081 | 1.078 | 1.075 | [14 empty] | 5 | 4.5 | 0.183 | 0.208 | 0.232 | 0.258 | 0.285
kVmeas at kVnom |  |  |  | 28.1 | 28.1 | 28.1 |  |  |  |  | 28 | 28.1 | [10 empty] | 6 | 6 | 1.08 | 1.078 | 1.074 | 1.074 | 1.071 |  |  | 6 | 6 | 1.164 | 1.16 | 1.151 | 1.15 | 1.144 |  |  |  | Factor de retrodispersión | [10 empty] | 6 | 5 | 0.164 | 0.187 | 0.209 | 0.232 | 0.258
Output at kVnom |  |  |  | 29.6 | 24.6 | 23.5 |  |  |  |  | 29 | 29.3 | [10 empty] | 7 | 7 | 1.152 | 1.147 | 1.141 | 1.138 | 1.135 |  |  | 7 | 7 | 1.232 | 1.225 | 1.214 | 1.208 | 1.204 | [14 empty] | 7 | 6 | 0.135 | 0.154 | 0.172 | 0.192 | 0.214
HVL at kVnom |  |  |  | 0.36 | 0.44 | 0.45 |  |  |  |  | 30 | 30.2 | [10 empty] | 8 | 8 | 1.22 | 1.213 | 1.206 | 1.205 | 1.199 |  |  | 8 | 8 | 1.275 | 1.265 | 1.257 | 1.254 | 1.247 | [7 empty] | HVL | HVL 1 |  | HVL 2 | fr1 | fr2 | fr | 8 | 7 | 0.114 | 0.13 | 0.145 | 0.163 | 0.177
[11 empty] | 31 | 31.2 | [10 empty] | 9 | 9 | 1.27 | 1.264 | 1.254 | 1.248 | 1.244 |  |  | 9 | 9 | 1.299 | 1.292 | 1.282 | 1.275 | 1.27 |  |  |  |  | HVL | fr |  | Indice | 4 |  | 5 |  |  |  | 9 | 8 | 0.098 | 0.112 | 0.126 | 0.14 | 0.154
Focus-Image distance (cm): |  |  |  |  | 66 | [5 empty] | 32 | 32.2 | [10 empty] | 10 | 10 | 1.295 | 1.287 | 1.279 | 1.275 | 1.272 |  |  | 10 | 10 | 1.307 | 1.298 | 1.29 | 1.286 | 1.283 |  |  |  | 1 | 0.25 | 1.07 | 1 | 0.41744 | 0.4 |  | 0.45 | 1.09 | 1.1 | 1.0935 | 10 | 9 | 0.0859 | 0.0981 | 0.1106 | 0.1233 | 0.1357
Platform- Image distance (cm): |  |  |  |  | 2 | [5 empty] | 33 | 33.1 | [10 empty] | 11 | 11 | 1.294 | 1.29 | 1.283 | 1.281 | 1.273 |  |  | 11 | 11 | 1.306 | 1.301 | 1.294 | 1.291 | 1.283 |  |  |  | 2 | 0.3 | 1.07 |  | Indice | 4 |  | 5 |  |  |  | 11 | 10 | 0.0763 | 0.0873 | 0.0986 | 0.1096 | 0.1207
Thickness crrection (cm): |  |  |  |  | 0.5 | [5 empty] | 34 | 34 | [29 empty] | 3 | 0.35 | 1.08 | 2 | 0.41744 | 0.4 |  | 0.45 | 1.09 | 1.1 | 1.0935 | 12 | 11 | 0.0687 | 0.0786 | 0.0887 | 0.0988 | 0.1088
[42 empty] | 4 | 0.4 | 1.09 |  | Indice | 3 |  | 4
|  |  |  | ESD = | K·fr |  | MGD = | K·g·s·c | [33 empty] | 5 | 0.45 | 1.1 | 3 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol.
[42 empty] | 6 | 0.5 | 1.11 |  | Indice | 2 |  | 3 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
Patient | Age | Thickness (cm) | kVps | mAs | Target/Filter |  |  | ESAK (mGy) | ESD (mGy) | MGD (mGy) |  |  | HVL | Espesor corregido | KVpm | DFPiel (cm) | Rend. 1m (mGy/mAs) | Factor c | Factor g | Factor s | Fr |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol. |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol. |  |  |  | 4 | 0.34973 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0799 | 1 | 5.5 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 0.21702 | 0.17898 | 0.198
1 | 45 | 5 | 26 | 162 | Mo-Rh |  |  | 8.9417 | 9.7777 | 1.8714 |  |  | 0.41744 | 5.5 | 25.9 | 58.5 | 18.8894 | 1.0393 | 0.198 | 1.017 | 1.0935 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [7 empty] | Indice | 2 |  | 3 | [5 empty] | 0.209 | 0.172 |  | 0.232 | 0.192
2 | 45 | 5.1 | 26 | 157 | Mo-Rh |  |  | 8.6954 | 9.5084 | 1.7968 |  |  | 0.41744 | 5.6 | 25.9 | 58.4 | 18.8894 | 1.0463 | 0.1942 | 1.017 | 1.0935 | 1 | 5.5 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 1.0047 | 1.074 | 1.0393 | 1 | 5.5 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 1.08 | 1.1507 | 1.1153 |  |  |  | 5 | 0.34973 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0799 |  | Indices | 6 | 7 |  | 4 | 5
3 | 44 | 4.1 | 28 | 62 | Mo-Mo |  |  | 5.2013 | 5.6269 | 1.0611 |  |  | 0.3591 | 4.6 | 28.1 | 59.4 | 29.6 | 0.98035 | 0.20809 | 1 | 1.0818 |  |  | 1.005 | 1.074 |  | 1.004 | 1.074 | [5 empty] | 1.081 | 1.151 |  | 1.078 | 1.15 | [7 empty] | Indice | 2 |  | 3 |  |  |  | 2 | 5.6 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 0.21702 | 0.17898 | 0.1942
4 | 54 | 4.9 | 27 | 114 | Mo-Mo |  |  | 8.6963 | 9.3915 | 1.6814 |  |  | 0.34973 | 5.4 | 27 | 58.6 | 26.1953 | 1.1132 | 0.17368 | 1 | 1.0799 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 6 | 0.33971 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0779 |  |  | 0.209 | 0.172 |  | 0.232 | 0.192
5 | 50 | 5.2 | 27 | 94 | Mo-Mo |  |  | 7.2446 | 7.8238 | 1.3487 |  |  | 0.34973 | 5.7 | 27 | 58.3 | 26.1953 | 1.1366 | 0.16379 | 1 | 1.0799 | 2 | 5.6 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 1.0047 | 1.074 | 1.0463 | 2 | 5.6 | 5 | 6 | 0.41744 | 0.4 | 0.45 | 1.08 | 1.1507 | 1.1224 |  |  |  |  | Indice | 3 |  | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4
6 | 45 | 3.9 | 26 | 84 | Mo-Mo |  |  | 5.4542 | 5.8794 | 1.0982 |  |  | 0.33971 | 4.4 | 25.9 | 59.6 | 23.0647 | 0.96743 | 0.20813 | 1 | 1.0779 |  |  | 1.005 | 1.074 |  | 1.004 | 1.074 | [5 empty] | 1.081 | 1.151 |  | 1.078 | 1.15 | [6 empty] | 7 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 | 3 | 4.6 | 4.5 | 5 | 0.3591 | 0.35 | 0.4 | 0.21237 | 0.191 | 0.20809
7 | 45 | 4.1 | 28 | 66 | Mo-Mo |  |  | 5.5368 | 5.9899 | 1.1295 |  |  | 0.3591 | 4.6 | 28.1 | 59.4 | 29.6 | 0.98035 | 0.20809 | 1 | 1.0818 |  | Indices | 4 | 5 |  | 3 | 4 |  |  |  |  | Indices | 4 | 5 |  | 3 | 4 | [7 empty] | Indice | 4 |  | 5 | [5 empty] | 0.208 | 0.187 |  | 0.232 | 0.209
8 | 43 | 4.4 | 28 | 103 | Rh-Rh |  |  | 6.93 | 7.6211 | 1.7364 |  |  | 0.44863 | 4.9 | 28.1 | 59.1 | 23.5 | 0.99832 | 0.23656 | 1.061 | 1.0997 | 3 | 4.6 | 4 | 5 | 0.3591 | 0.35 | 0.4 | 0.94336 | 1.005 | 0.98035 | 3 | 4.6 | 4 | 5 | 0.3591 | 0.35 | 0.4 | 1 | 1.0818 | 1.0491 |  |  |  | 8 | 0.44863 | 0.4 |  | 0.45 | 1.09 | 1.1 | 1.0997 |  | Indices | 6 | 7 |  | 2 | 3
9 | 43 | 4.8 | 28 | 126 | Rh-Rh |  |  | 8.5933 | 9.4503 | 2.0504 |  |  | 0.44863 | 5.3 | 28.1 | 58.7 | 23.5 | 1.025 | 0.2194 | 1.061 | 1.0997 |  |  | 0.943 | 1.005 |  | 0.945 | 1.005 | [5 empty] | 1 | 1.082 |  | 1 | 1.081 | [7 empty] | Indice | 4 |  | 5 |  |  |  | 4 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 0.18688 | 0.1539 | 0.17368
10 | 58 | 4.9 | 27 | 114 | Mo-Mo |  |  | 8.6963 | 9.3915 | 1.6814 |  |  | 0.34973 | 5.4 | 27 | 58.6 | 26.1953 | 1.1132 | 0.17368 | 1 | 1.0799 |  | Indices | 5 | 6 |  | 2 | 3 |  |  |  |  | Indices | 5 | 6 |  | 2 | 3 | [6 empty] | 9 | 0.44863 | 0.4 |  | 0.45 | 1.09 | 1.1 | 1.0997 |  |  | 0.164 | 0.135 |  | 0.187 | 0.154
11 | 58 | 4.9 | 27 | 107 | Mo-Mo |  |  | 8.1623 | 8.8148 | 1.5782 |  |  | 0.34973 | 5.4 | 27 | 58.6 | 26.1953 | 1.1132 | 0.17368 | 1 | 1.0799 | 4 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.005 | 1.078 | 1.0342 | 4 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.082 | 1.16 | 1.1132 |  |  |  |  | Indice | 2 |  | 3 |  |  |  |  | Indices | 6 | 7 |  | 2 | 3
12 | 49 | 5.3 | 28 | 116 | Mo-Mo |  |  | 10.1369 | 10.9663 | 1.7671 |  |  | 0.3591 | 5.8 | 28.1 | 58.2 | 29.6 | 1.0628 | 0.16402 | 1 | 1.0818 |  |  | 1.005 | 1.08 |  | 1.005 | 1.078 | [5 empty] | 1.086 | 1.164 |  | 1.082 | 1.16 | [6 empty] | 10 | 0.34973 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0799 | 5 | 5.7 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 0.18688 | 0.1539 | 0.16379
13 | 49 | 5.6 | 28 | 126 | Mo-Mo |  |  | 11.1251 | 12.0354 | 1.8673 |  |  | 0.3591 | 6.1 | 28.1 | 57.9 | 29.6 | 1.0841 | 0.15482 | 1 | 1.0818 |  | Indices | 5 | 6 |  | 2 | 3 |  |  |  |  | Indices | 5 | 6 |  | 2 | 3 | [7 empty] | Indice | 2 |  | 3 | [5 empty] | 0.164 | 0.135 |  | 0.187 | 0.154
14 | 56 | 5.1 | 28 | 112 | Mo-Mo |  |  | 9.7204 | 10.5157 | 1.872 |  |  | 0.3591 | 5.6 | 28.1 | 58.4 | 29.6 | 1.1277 | 0.17077 | 1 | 1.0818 | 5 | 5.7 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.005 | 1.078 | 1.0561 | 5 | 5.7 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.082 | 1.16 | 1.1366 |  |  |  | 11 | 0.34973 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0799 |  | Indices | 4 | 5 |  | 2 | 3
15 | 56 | 5.5 | 28 | 134 | Mo-Mo |  |  | 11.7907 | 12.7554 | 2.1481 |  |  | 0.3591 | 6 | 28.1 | 58 | 29.6 | 1.1584 | 0.15728 | 1 | 1.0818 |  |  | 1.005 | 1.08 |  | 1.005 | 1.078 | [5 empty] | 1.086 | 1.164 |  | 1.082 | 1.16 | [7 empty] | Indice | 3 |  | 4 |  |  |  | 6 | 4.4 | 4 | 4.5 | 0.33971 | 0.3 | 0.35 | 0.22924 | 0.20285 | 0.20813
16 | 58 | 4.9 | 27 | 107 | Mo-Mo |  |  | 8.1623 | 8.8148 | 1.5782 |  |  | 0.34973 | 5.4 | 27 | 58.6 | 26.1953 | 1.1132 | 0.17368 | 1 | 1.0799 |  | Indices | 4 | 5 |  | 2 | 3 |  |  |  |  | Indices | 4 | 5 |  | 2 | 3 | [6 empty] | 12 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  |  | 0.207 | 0.183 |  | 0.235 | 0.208
17 | 49 | 5.3 | 28 | 116 | Mo-Mo |  |  | 10.1369 | 10.9663 | 1.7671 |  |  | 0.3591 | 5.8 | 28.1 | 58.2 | 29.6 | 1.0628 | 0.16402 | 1 | 1.0818 | 6 | 4.4 | 4 | 5 | 0.33971 | 0.3 | 0.35 | 0.94238 | 1.005 | 0.96743 | 6 | 4.4 | 4 | 5 | 0.33971 | 0.3 | 0.35 | 1 | 1.0828 | 1.0331 |  |  |  |  | Indice | 3 |  | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4
18 | 49 | 5.6 | 28 | 126 | Mo-Mo |  |  | 11.1251 | 12.0354 | 1.8673 |  |  | 0.3591 | 6.1 | 28.1 | 57.9 | 29.6 | 1.0841 | 0.15482 | 1 | 1.0818 |  |  | 0.94 | 1.005 |  | 0.943 | 1.005 | [5 empty] | 1 | 1.086 |  | 1 | 1.082 | [6 empty] | 13 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 | 7 | 4.6 | 4.5 | 5 | 0.3591 | 0.35 | 0.4 | 0.21237 | 0.191 | 0.20809
19 | 56 | 5.1 | 28 | 112 | Mo-Mo |  |  | 9.7204 | 10.5157 | 1.872 |  |  | 0.3591 | 5.6 | 28.1 | 58.4 | 29.6 | 1.1277 | 0.17077 | 1 | 1.0818 |  | Indices | 4 | 5 |  | 3 | 4 |  |  |  |  | Indices | 4 | 5 |  | 3 | 4 | [7 empty] | Indice | 3 |  | 4 | [5 empty] | 0.208 | 0.187 |  | 0.232 | 0.209
20 | 56 | 5.1 | 28 | 113 | Mo-Mo |  |  | 9.8072 | 10.6096 | 1.8887 |  |  | 0.3591 | 5.6 | 28.1 | 58.4 | 29.6 | 1.1277 | 0.17077 | 1 | 1.0818 | 7 | 4.6 | 4 | 5 | 0.3591 | 0.35 | 0.4 | 0.94336 | 1.005 | 0.98035 | 7 | 4.6 | 4 | 5 | 0.3591 | 0.35 | 0.4 | 1 | 1.0818 | 1.0491 |  |  |  | 14 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  | Indices | 5 | 6 |  | 4 | 5
|  | Mean | 27.45 | 112.55 | [19 empty] | 0.943 | 1.005 |  | 0.945 | 1.005 | [5 empty] | 1 | 1.082 |  | 1 | 1.081 | [7 empty] | Indice | 3 |  | 4 |  |  |  | 8 | 4.9 | 4.5 | 5 | 0.44863 | 0.4 | 0.45 | 0.25729 | 0.23137 | 0.23656
[23 empty] | Indices | 4 | 5 |  | 4 | 5 |  |  |  |  | Indices | 4 | 5 |  | 4 | 5 | [6 empty] | 15 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  |  | 0.232 | 0.209 |  | 0.258 | 0.232
[5 empty] | ESAK (mGy) | ESD (mGy) | MGD (mGy) | Corrected breast thickness (cm) | [13 empty] | 8 | 4.9 | 4 | 5 | 0.44863 | 0.4 | 0.45 | 0.94695 | 1.004 | 0.99832 | 8 | 4.9 | 4 | 5 | 0.44863 | 0.4 | 0.45 | 1 | 1.0781 | 1.0703 |  |  |  |  | Indice | 2 |  | 3 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
|  |  |  | Mean | 8.6939 | 9.4245 | 1.6831 | 5.44 | [15 empty] | 0.945 | 1.005 |  | 0.947 | 1.004 | [5 empty] | 1 | 1.081 |  | 1 | 1.078 | [6 empty] | 16 | 0.34973 | 0.3 |  | 0.35 | 1.07 | 1.08 | 1.0799 | 9 | 5.3 | 5 | 6 | 0.44863 | 0.4 | 0.45 | 0.23137 | 0.19145 | 0.2194
|  |  |  | CV (%) | 21.6693 | 21.6016 | 18.1201 | 8.8742 | [14 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 1 | Indice | 3 |  | 4 | [5 empty] | 0.209 | 0.172 |  | 0.232 | 0.192
[22 empty] | 9 | 5.3 | 5 | 6 | 0.44863 | 0.4 | 0.45 | 1.004 | 1.074 | 1.025 | 9 | 5.3 | 5 | 6 | 0.44863 | 0.4 | 0.45 | 1.0781 | 1.15 | 1.0997 |  |  |  | 17 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  | Indices | 6 | 7 |  | 2 | 3
[24 empty] | 1.005 | 1.074 |  | 1.004 | 1.074 | [5 empty] | 1.081 | 1.151 |  | 1.078 | 1.15 | [7 empty] | Indice | 3 |  | 4 |  |  |  | 10 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 0.18688 | 0.1539 | 0.17368
[23 empty] | Indices | 5 | 6 |  | 2 | 3 |  |  |  |  | Indices | 5 | 6 |  | 2 | 3 | [6 empty] | 18 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  |  | 0.164 | 0.135 |  | 0.187 | 0.154
[22 empty] | 10 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.005 | 1.078 | 1.0342 | 10 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.082 | 1.16 | 1.1132 |  |  |  |  | Indice | 3 |  | 4 |  |  |  |  | Indices | 6 | 7 |  | 2 | 3
[24 empty] | 1.005 | 1.08 |  | 1.005 | 1.078 | [5 empty] | 1.086 | 1.164 |  | 1.082 | 1.16 | [6 empty] | 19 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 | 11 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 0.18688 | 0.1539 | 0.17368
[23 empty] | Indices | 5 | 6 |  | 2 | 3 |  |  |  |  | Indices | 5 | 6 |  | 2 | 3 | [7 empty] | Indice | 3 |  | 4 | [5 empty] | 0.164 | 0.135 |  | 0.187 | 0.154
[22 empty] | 11 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.005 | 1.078 | 1.0342 | 11 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.082 | 1.16 | 1.1132 |  |  |  | 20 | 0.3591 | 0.35 |  | 0.4 | 1.08 | 1.09 | 1.0818 |  | Indices | 6 | 7 |  | 3 | 4
[24 empty] | 1.005 | 1.08 |  | 1.005 | 1.078 | [5 empty] | 1.086 | 1.164 |  | 1.082 | 1.16 | [14 empty] | 12 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 0.191 | 0.15728 | 0.16402
[23 empty] | Indices | 5 | 6 |  | 3 | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4 | [16 empty] | 0.187 | 0.154 |  | 0.209 | 0.172
[22 empty] | 12 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.005 | 1.0773 | 1.0628 | 12 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.0818 | 1.1584 | 1.1431 | [12 empty] | Indices | 7 | 8 |  | 3 | 4
[24 empty] | 1.005 | 1.078 |  | 1.005 | 1.074 | [5 empty] | 1.082 | 1.16 |  | 1.081 | 1.151 | [14 empty] | 13 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 0.15728 | 0.13273 | 0.15482
[23 empty] | Indices | 6 | 7 |  | 3 | 4 |  |  |  |  | Indices | 6 | 7 |  | 3 | 4 | [16 empty] | 0.154 | 0.13 |  | 0.172 | 0.145
[22 empty] | 13 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.0773 | 1.1459 | 1.0841 | 13 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.1584 | 1.223 | 1.1648 | [12 empty] | Indices | 6 | 7 |  | 3 | 4
[24 empty] | 1.078 | 1.147 |  | 1.074 | 1.141 | [5 empty] | 1.16 | 1.225 |  | 1.151 | 1.214 | [14 empty] | 14 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 0.191 | 0.15728 | 0.17077
[23 empty] | Indices | 5 | 6 |  | 3 | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4 | [16 empty] | 0.187 | 0.154 |  | 0.209 | 0.172
[22 empty] | 14 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.005 | 1.0773 | 1.0484 | 14 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.0818 | 1.1584 | 1.1277 | [12 empty] | Indices | 7 | 8 |  | 3 | 4
[24 empty] | 1.005 | 1.078 |  | 1.005 | 1.074 | [5 empty] | 1.082 | 1.16 |  | 1.081 | 1.151 | [14 empty] | 15 | 6 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 0.15728 | 0.13273 | 0.15728
[23 empty] | Indices | 6 | 7 |  | 3 | 4 |  |  |  |  | Indices | 6 | 7 |  | 3 | 4 | [16 empty] | 0.154 | 0.13 |  | 0.172 | 0.145
[22 empty] | 15 | 6 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.0773 | 1.1459 | 1.0773 | 15 | 6 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.1584 | 1.223 | 1.1584 | [12 empty] | Indices | 6 | 7 |  | 2 | 3
[24 empty] | 1.078 | 1.147 |  | 1.074 | 1.141 | [5 empty] | 1.16 | 1.225 |  | 1.151 | 1.214 | [14 empty] | 16 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 0.18688 | 0.1539 | 0.17368
[23 empty] | Indices | 5 | 6 |  | 2 | 3 |  |  |  |  | Indices | 5 | 6 |  | 2 | 3 | [16 empty] | 0.164 | 0.135 |  | 0.187 | 0.154
[22 empty] | 16 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.005 | 1.078 | 1.0342 | 16 | 5.4 | 5 | 6 | 0.34973 | 0.3 | 0.35 | 1.082 | 1.16 | 1.1132 | [12 empty] | Indices | 6 | 7 |  | 3 | 4
[24 empty] | 1.005 | 1.08 |  | 1.005 | 1.078 | [5 empty] | 1.086 | 1.164 |  | 1.082 | 1.16 | [14 empty] | 17 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 0.191 | 0.15728 | 0.16402
[23 empty] | Indices | 5 | 6 |  | 3 | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4 | [16 empty] | 0.187 | 0.154 |  | 0.209 | 0.172
[22 empty] | 17 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.005 | 1.0773 | 1.0628 | 17 | 5.8 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.0818 | 1.1584 | 1.1431 | [12 empty] | Indices | 7 | 8 |  | 3 | 4
[24 empty] | 1.005 | 1.078 |  | 1.005 | 1.074 | [5 empty] | 1.082 | 1.16 |  | 1.081 | 1.151 | [14 empty] | 18 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 0.15728 | 0.13273 | 0.15482
[23 empty] | Indices | 6 | 7 |  | 3 | 4 |  |  |  |  | Indices | 6 | 7 |  | 3 | 4 | [16 empty] | 0.154 | 0.13 |  | 0.172 | 0.145
[22 empty] | 18 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.0773 | 1.1459 | 1.0841 | 18 | 6.1 | 6 | 7 | 0.3591 | 0.35 | 0.4 | 1.1584 | 1.223 | 1.1648 | [12 empty] | Indices | 6 | 7 |  | 3 | 4
[24 empty] | 1.078 | 1.147 |  | 1.074 | 1.141 | [5 empty] | 1.16 | 1.225 |  | 1.151 | 1.214 | [14 empty] | 19 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 0.191 | 0.15728 | 0.17077
[23 empty] | Indices | 5 | 6 |  | 3 | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4 | [16 empty] | 0.187 | 0.154 |  | 0.209 | 0.172
[22 empty] | 19 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.005 | 1.0773 | 1.0484 | 19 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.0818 | 1.1584 | 1.1277 | [12 empty] | Indices | 6 | 7 |  | 3 | 4
[24 empty] | 1.005 | 1.078 |  | 1.005 | 1.074 | [5 empty] | 1.082 | 1.16 |  | 1.081 | 1.151 | [14 empty] | 20 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 0.191 | 0.15728 | 0.17077
[23 empty] | Indices | 5 | 6 |  | 3 | 4 |  |  |  |  | Indices | 5 | 6 |  | 3 | 4 | [16 empty] | 0.187 | 0.154 |  | 0.209 | 0.172
[22 empty] | 20 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.005 | 1.0773 | 1.0484 | 20 | 5.6 | 5 | 6 | 0.3591 | 0.35 | 0.4 | 1.0818 | 1.1584 | 1.1277
[24 empty] | 1.005 | 1.078 |  | 1.005 | 1.074 | [5 empty] | 1.082 | 1.16 |  | 1.081 | 1.151
Note on B17: (mm Al) measured behind compression plate
Note on B21: fpro: Difference between real and measured compressed breast thickness
Note on D25: Measured or showed on display
Note on E7: Fill only cells in blue color
Note on G21: Real thickness-thickness showed on display
Note on G23: fisica1: fr is HVL dependent
Note on G25: Mo-Mo Mo-Rh Rh-Rh

## CalMammo_2
Factor s
Patient dose estimation in mammography | [23 empty] | W-Rh | 1.042
[24 empty] | Rh-Al | 1.044
Craniocaudal Projection
Date | [23 empty] | Factor c grupo edad 40 a 49 años | [8 empty] | Factor c grupo edad 50 a 64 años | [20 empty] | Valor de g
[24 empty] | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | [15 empty] | 1 | 2 | 3 | 4 | 5 | 6
Facility: | [10 empty] | kVp set | kVpmeas | [11 empty] | Espesor | HVL (mm Al) | [7 empty] | Espesor | HVL (mm Al) | [7 empty] | Combinación ánodo/filtro |  | Espesor de filtro |  | n | a | b | A | c |  |  |  | Espesor | HVL (mm Al)
Room: | [10 empty] | 23 | 23 | [10 empty] | 1 | mama | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  |  | 1 | mama | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  |  |  | W/50 mm Rh |  | 58,9 mm |  | 1.96 | -0.00054 | 0.0403 | 0.042838 | -0.16604 | W-Rh |  | 1 | mama | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
Unit: | [10 empty] | 24 | 24 | [10 empty] | 2 | 2 | 0.9 | 0.905 | 0.91 | 0.914 | 0.919 |  |  | 2 | 2 | 0.9 | 0.905 | 0.91 | 0.914 | 0.919 |  |  |  | Rh/1,0 mm Al |  | 1,2 mm |  | 4.39 | -0.00113 | 0.0909 | 1.0743e-05 | -1.102 | Rh-Al |  | 2 | 2 | 0.473 | 0.509 | 0.543 | 0.573 | 0.587
[11 empty] | 25 | 24.9 | [10 empty] | 3 | 3 | 0.903 | 0.906 | 0.911 | 0.915 | 0.918 |  |  | 3 | 3 | 0.931 | 0.933 | 0.937 | 0.94 | 0.741 |  |  |  | Mo/1,0 mm Al |  | 1,2 mm |  | 4.23 | -0.000775 | 0.0593 | #DIV/0! | -1.0544 | Mo-Al |  | 3 | 3 | 0.342 | 0.374 | 0.406 | 0.437 | 0.466
|  |  |  | W-Rh | Rh-Al | [5 empty] | 26 | 25.9 | [10 empty] | 4 | 4 | 0.945 | 0.947 | 0.948 | 0.952 | 0.955 |  |  | 4 | 4 | 1 | 1 | 1 | 1 | 1 | [14 empty] | 4 | 4 | 0.261 | 0.289 | 0.318 | 0.346 | 0.374
kV ( nominal) |  |  |  | 28 | 28 | [5 empty] | 27 | 27 | [10 empty] | 5 | 5 | 1.005 | 1.004 | 1.004 | 1.004 | 1.004 |  |  | 5 | 5 | 1.081 | 1.078 | 1.075 | 1.071 | 1.069 | [14 empty] | 5 | 4.5 | 0.232 | 0.258 | 0.285 | 0.311 | 0.339
kVmeas at kVnom |  |  |  | 28.1 | 28.1 | [5 empty] | 28 | 28.1 | [10 empty] | 6 | 6 | 1.074 | 1.074 | 1.071 | 1.068 | 1.066 |  |  | 6 | 6 | 1.151 | 1.15 | 1.144 | 1.139 | 1.134 |  |  |  | Factor de retrodispersión | [10 empty] | 6 | 5 | 0.209 | 0.232 | 0.258 | 0.287 | 0.31
Output at kVnom |  |  |  | 29.6 | 24.6 | [5 empty] | 29 | 29.3 | [10 empty] | 7 | 7 | 1.141 | 1.138 | 1.135 | 1.13 | 1.127 |  |  | 7 | 7 | 1.214 | 1.208 | 1.204 | 1.196 | 1.188 | [14 empty] | 7 | 6 | 0.172 | 0.192 | 0.214 | 0.236 | 0.261
HVL at kVnom |  |  |  | 0.54 | 0.56 | [5 empty] | 30 | 30.2 | [10 empty] | 8 | 8 | 1.206 | 1.205 | 1.199 | 1.19 | 1.183 |  |  | 8 | 8 | 1.257 | 1.254 | 1.247 | 1.237 | 1.227 | [7 empty] | HVL | HVL 1 |  | HVL 2 | fr1 | fr2 | fr | 8 | 7 | 0.145 | 0.163 | 0.177 | 0.202 | 0.224
[11 empty] | 31 | 31.2 | [10 empty] | 9 | 9 | 1.254 | 1.248 | 1.244 | 1.235 | 1.225 |  |  | 9 | 9 | 1.282 | 1.275 | 1.27 | 1.26 | 1.249 |  |  |  |  | HVL | fr |  | Indice | 5 |  | 6 |  |  |  | 9 | 8 | 0.126 | 0.14 | 0.154 | 0.175 | 0.195
Focus-Image distance (cm): |  |  |  |  | 66 | [5 empty] | 32 | 32.2 | [10 empty] | 10 | 10 | 1.279 | 1.275 | 1.272 | 1.262 | 1.251 |  |  | 10 | 10 | 1.29 | 1.286 | 1.283 | 1.272 | 1.261 |  |  |  | 1 | 0.25 | 1.07 | 1 | 0.51672 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 10 | 9 | 0.1106 | 0.1233 | 0.1357 | 0.1543 | 0.1723
Platform- Image distance (cm): |  |  |  |  | 2 | [5 empty] | 33 | 33.1 | [10 empty] | 11 | 11 | 1.283 | 1.281 | 1.273 | 1.264 | 1.256 |  |  | 11 | 11 | 1.294 | 1.291 | 1.283 | 1.274 | 1.266 |  |  |  | 2 | 0.35 | 1.08 |  | Indice | 5 |  | 6 |  |  |  | 11 | 10 | 0.0986 | 0.1096 | 0.1207 | 0.1375 | 0.154
Thickness crrection (cm): |  |  |  |  | 0.5 | [5 empty] | 34 | 34 | [29 empty] | 3 | 0.4 | 1.09 | 2 | 0.51672 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 12 | 11 | 0.0887 | 0.0988 | 0.1088 | 0.124 | 0.1385
[42 empty] | 4 | 0.45 | 1.1 |  | Indice | 5 |  | 6
|  |  |  | ESD = | K·fr |  | MGD = | K·g·s·c | [33 empty] | 5 | 0.5 | 1.11 | 3 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol.
[42 empty] | 6 | 0.6 | 1.12 |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
Patient | Age | Thickness (cm) | kVps | mAs | Target/Filter |  |  | ESAK (mGy) | ESD (mGy) | MGD (mGy) |  |  | HVL | Espesor corregido | KVpm | DFPiel (cm) | Rend. 1m (mGy/mAs) | Factor c | Factor g | Factor s | Fr |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol. |  | E correg | E1 | E2 | HVL | HVL1 | HVL2 | Int.1 | Int.2 | Valor interpol. |  |  |  | 4 | 0.5285 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 1 | 5.5 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 0.258 | 0.214 | 0.236
1 | 45 | 5 | 26 | 162 | W-Rh |  |  | 11.9426 | 13.2563 | 3.047 |  |  | 0.51672 | 5.5 | 25.9 | 58.5 | 25.2287 | 1.0375 | 0.236 | 1.042 | 1.11 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [7 empty] | Indice | 5 |  | 6 | [5 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
2 | 45 | 5.1 | 26 | 157 | W-Rh |  |  | 11.6137 | 12.8912 | 2.9266 |  |  | 0.51672 | 5.6 | 25.9 | 58.4 | 25.2287 | 1.0442 | 0.2316 | 1.042 | 1.11 | 1 | 5.5 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0375 | 1 | 5.5 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1095 |  |  |  | 5 | 0.5285 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 6 | 7 |  | 4 | 5
3 | 44 | 4.1 | 28 | 62 | W-Rh |  |  | 5.2013 | 5.7734 | 1.4875 |  |  | 0.539 | 4.6 | 28.1 | 59.4 | 29.6 | 0.9816 | 0.2796 | 1.042 | 1.11 |  |  | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [7 empty] | Indice | 4 |  | 5 |  |  |  | 2 | 5.6 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2316
4 | 54 | 4.9 | 27 | 114 | Rh-Al |  |  | 6.8535 | 7.6074 | 1.8965 |  |  | 0.5285 | 5.4 | 27 | 58.6 | 20.6444 | 1.1026 | 0.2404 | 1.044 | 1.11 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 6 | 0.49749 | 0.45 |  | 0.5 | 1.1 | 1.11 | 1.1095 |  |  | 0.258 | 0.214 |  | 0.258 | 0.214
5 | 50 | 5.2 | 27 | 94 | Rh-Al |  |  | 5.7094 | 6.3375 | 1.5212 |  |  | 0.5285 | 5.7 | 27 | 58.3 | 20.6444 | 1.1233 | 0.2272 | 1.044 | 1.11 | 2 | 5.6 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0442 | 2 | 5.6 | 5 | 6 | 0.51672 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1164 |  |  |  |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5
6 | 45 | 3.9 | 26 | 84 | Rh-Al |  |  | 4.0671 | 4.5125 | 1.1958 |  |  | 0.49749 | 4.4 | 25.9 | 59.6 | 17.1989 | 0.97037 | 0.29022 | 1.044 | 1.1095 |  |  | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [6 empty] | 7 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 3 | 4.6 | 4.5 | 5 | 0.539 | 0.5 | 0.55 | 0.285 | 0.258 | 0.2796
7 | 45 | 4.1 | 28 | 66 | Rh-Al |  |  | 4.6016 | 5.1077 | 1.3185 |  |  | 0.55725 | 4.6 | 28.1 | 59.4 | 24.6 | 0.9816 | 0.2796 | 1.044 | 1.11 |  | Indices | 4 | 5 |  | 4 | 5 |  |  |  |  | Indices | 4 | 5 |  | 4 | 5 | [7 empty] | Indice | 5 |  | 6 | [5 empty] | 0.285 | 0.258 |  | 0.285 | 0.258
8 | 43 | 4.4 | 28 | 103 | Rh-Al |  |  | 7.2543 | 8.0523 | 1.9917 |  |  | 0.55725 | 4.9 | 28.1 | 59.1 | 24.6 | 0.9984 | 0.2634 | 1.044 | 1.11 | 3 | 4.6 | 4 | 5 | 0.539 | 0.5 | 0.55 | 0.948 | 1.004 | 0.9816 | 3 | 4.6 | 4 | 5 | 0.539 | 0.5 | 0.55 | 1 | 1.075 | 1.045 |  |  |  | 8 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 6 | 7 |  | 4 | 5
9 | 43 | 4.8 | 28 | 126 | Rh-Al |  |  | 8.9956 | 9.9851 | 2.3544 |  |  | 0.55725 | 5.3 | 28.1 | 58.7 | 24.6 | 1.0241 | 0.2448 | 1.044 | 1.11 |  |  | 0.948 | 1.004 |  | 0.948 | 1.004 | [5 empty] | 1 | 1.075 |  | 1 | 1.075 | [7 empty] | Indice | 5 |  | 6 |  |  |  | 4 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2404
10 | 58 | 4.9 | 27 | 114 | Rh-Al |  |  | 6.8535 | 7.6074 | 1.8965 |  |  | 0.5285 | 5.4 | 27 | 58.6 | 20.6444 | 1.1026 | 0.2404 | 1.044 | 1.11 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 9 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  |  | 0.258 | 0.214 |  | 0.258 | 0.214
11 | 58 | 4.9 | 27 | 107 | Rh-Al |  |  | 6.4326 | 7.1402 | 1.7801 |  |  | 0.5285 | 5.4 | 27 | 58.6 | 20.6444 | 1.1026 | 0.2404 | 1.044 | 1.11 | 4 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0308 | 4 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1026 |  |  |  |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
12 | 49 | 5.3 | 28 | 116 | Rh-Al |  |  | 8.4246 | 9.3513 | 2.0725 |  |  | 0.55725 | 5.8 | 28.1 | 58.2 | 24.6 | 1.0576 | 0.2228 | 1.044 | 1.11 |  |  | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [6 empty] | 10 | 0.5285 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 5 | 5.7 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2272
13 | 49 | 5.6 | 28 | 126 | Rh-Al |  |  | 9.2459 | 10.2629 | 2.1871 |  |  | 0.55725 | 6.1 | 28.1 | 57.9 | 24.6 | 1.0774 | 0.2103 | 1.044 | 1.11 |  | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [7 empty] | Indice | 5 |  | 6 | [5 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
14 | 56 | 5.1 | 28 | 112 | Rh-Al |  |  | 8.0784 | 8.9671 | 2.1807 |  |  | 0.55725 | 5.6 | 28.1 | 58.4 | 24.6 | 1.1164 | 0.2316 | 1.044 | 1.11 | 5 | 5.7 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0509 | 5 | 5.7 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1233 |  |  |  | 11 | 0.5285 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 4 | 5 |  | 3 | 4
15 | 56 | 5.5 | 28 | 134 | W-Rh |  |  | 11.7907 | 13.0877 | 3.0078 |  |  | 0.539 | 6 | 28.1 | 58 | 29.6 | 1.144 | 0.214 | 1.042 | 1.11 |  |  | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [7 empty] | Indice | 5 |  | 6 |  |  |  | 6 | 4.4 | 4 | 4.5 | 0.49749 | 0.45 | 0.5 | 0.31654 | 0.28364 | 0.29022
16 | 58 | 4.9 | 27 | 107 | W-Rh |  |  | 8.5288 | 9.467 | 2.3556 |  |  | 0.5284 | 5.4 | 27 | 58.6 | 27.3716 | 1.1026 | 0.2404 | 1.042 | 1.11 |  | Indices | 4 | 5 |  | 3 | 4 |  |  |  |  | Indices | 4 | 5 |  | 3 | 4 | [6 empty] | 12 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  |  | 0.289 | 0.258 |  | 0.318 | 0.285
17 | 49 | 5.3 | 28 | 116 | W-Rh |  |  | 10.1369 | 11.2519 | 2.4889 |  |  | 0.539 | 5.8 | 28.1 | 58.2 | 29.6 | 1.0576 | 0.2228 | 1.042 | 1.11 | 6 | 4.4 | 4 | 5 | 0.49749 | 0.45 | 0.5 | 0.94795 | 1.004 | 0.97037 | 6 | 4.4 | 4 | 5 | 0.49749 | 0.45 | 0.5 | 1 | 1.0752 | 1.0301 |  |  |  |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5
18 | 49 | 5.6 | 28 | 126 | W-Rh |  |  | 11.1251 | 12.3489 | 2.6266 |  |  | 0.539 | 6.1 | 28.1 | 57.9 | 29.6 | 1.0774 | 0.2103 | 1.042 | 1.11 |  |  | 0.947 | 1.004 |  | 0.948 | 1.004 | [5 empty] | 1 | 1.078 |  | 1 | 1.075 | [6 empty] | 13 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 7 | 4.6 | 4.5 | 5 | 0.55725 | 0.5 | 0.55 | 0.285 | 0.258 | 0.2796
19 | 56 | 5.1 | 28 | 112 | W-Rh |  |  | 9.7204 | 10.7896 | 2.6188 |  |  | 0.539 | 5.6 | 28.1 | 58.4 | 29.6 | 1.1164 | 0.2316 | 1.042 | 1.11 |  | Indices | 4 | 5 |  | 4 | 5 |  |  |  |  | Indices | 4 | 5 |  | 4 | 5 | [7 empty] | Indice | 5 |  | 6 | [5 empty] | 0.285 | 0.258 |  | 0.285 | 0.258
20 | 56 | 5.1 | 28 | 113 | W-Rh |  |  | 9.8072 | 10.886 | 2.6422 |  |  | 0.539 | 5.6 | 28.1 | 58.4 | 29.6 | 1.1164 | 0.2316 | 1.042 | 1.11 | 7 | 4.6 | 4 | 5 | 0.55725 | 0.5 | 0.55 | 0.948 | 1.004 | 0.9816 | 7 | 4.6 | 4 | 5 | 0.55725 | 0.5 | 0.55 | 1 | 1.075 | 1.045 |  |  |  | 14 | 0.55725 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 5 | 6 |  | 4 | 5
|  | Mean | 27.45 | 112.55 | [19 empty] | 0.948 | 1.004 |  | 0.948 | 1.004 | [5 empty] | 1 | 1.075 |  | 1 | 1.075 | [7 empty] | Indice | 5 |  | 6 |  |  |  | 8 | 4.9 | 4.5 | 5 | 0.55725 | 0.5 | 0.55 | 0.285 | 0.258 | 0.2634
[23 empty] | Indices | 4 | 5 |  | 4 | 5 |  |  |  |  | Indices | 4 | 5 |  | 4 | 5 | [6 empty] | 15 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  |  | 0.285 | 0.258 |  | 0.285 | 0.258
[5 empty] | ESAK (mGy) | ESD (mGy) | MGD (mGy) | Corrected breast thickness (cm) | [13 empty] | 8 | 4.9 | 4 | 5 | 0.55725 | 0.5 | 0.55 | 0.948 | 1.004 | 0.9984 | 8 | 4.9 | 4 | 5 | 0.55725 | 0.5 | 0.55 | 1 | 1.075 | 1.0675 |  |  |  |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
|  |  |  | Mean | 8.3192 | 9.2342 | 2.1798 | 5.44 | [15 empty] | 0.948 | 1.004 |  | 0.948 | 1.004 | [5 empty] | 1 | 1.075 |  | 1 | 1.075 | [6 empty] | 16 | 0.5284 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 9 | 5.3 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2448
|  |  |  | CV (%) | 28.9544 | 28.9568 | 25.2343 | 8.8742 | [14 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 1 | Indice | 5 |  | 6 | [5 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 9 | 5.3 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0241 | 9 | 5.3 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.075 | 1.144 | 1.0957 |  |  |  | 17 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [7 empty] | Indice | 5 |  | 6 |  |  |  | 10 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2404
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [6 empty] | 18 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  |  | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 10 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0308 | 10 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1026 |  |  |  |  | Indice | 5 |  | 6 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [6 empty] | 19 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 | 11 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2404
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [7 empty] | Indice | 5 |  | 6 | [5 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 11 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0308 | 11 | 5.4 | 5 | 6 | 0.5285 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1026 |  |  |  | 20 | 0.539 | 0.5 |  | 0.6 | 1.11 | 1.11 | 1.11 |  | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 12 | 5.8 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2228
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 12 | 5.8 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0576 | 12 | 5.8 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1302 | [12 empty] | Indices | 7 | 8 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 13 | 6.1 | 6 | 7 | 0.55725 | 0.5 | 0.55 | 0.214 | 0.177 | 0.2103
[23 empty] | Indices | 6 | 7 |  | 4 | 5 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5 | [16 empty] | 0.214 | 0.177 |  | 0.214 | 0.177
[22 empty] | 13 | 6.1 | 6 | 7 | 0.55725 | 0.5 | 0.55 | 1.071 | 1.135 | 1.0774 | 13 | 6.1 | 6 | 7 | 0.55725 | 0.5 | 0.55 | 1.144 | 1.204 | 1.15 | [12 empty] | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.071 | 1.135 |  | 1.071 | 1.135 | [5 empty] | 1.144 | 1.204 |  | 1.144 | 1.204 | [14 empty] | 14 | 5.6 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2316
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 14 | 5.6 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0442 | 14 | 5.6 | 5 | 6 | 0.55725 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1164 | [12 empty] | Indices | 7 | 8 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 15 | 6 | 6 | 7 | 0.539 | 0.5 | 0.55 | 0.214 | 0.177 | 0.214
[23 empty] | Indices | 6 | 7 |  | 4 | 5 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5 | [16 empty] | 0.214 | 0.177 |  | 0.214 | 0.177
[22 empty] | 15 | 6 | 6 | 7 | 0.539 | 0.5 | 0.55 | 1.071 | 1.135 | 1.071 | 15 | 6 | 6 | 7 | 0.539 | 0.5 | 0.55 | 1.144 | 1.204 | 1.144 | [12 empty] | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.071 | 1.135 |  | 1.071 | 1.135 | [5 empty] | 1.144 | 1.204 |  | 1.144 | 1.204 | [14 empty] | 16 | 5.4 | 5 | 6 | 0.5284 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2404
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 16 | 5.4 | 5 | 6 | 0.5284 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0308 | 16 | 5.4 | 5 | 6 | 0.5284 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1026 | [12 empty] | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 17 | 5.8 | 5 | 6 | 0.539 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2228
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 17 | 5.8 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0576 | 17 | 5.8 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1302 | [12 empty] | Indices | 7 | 8 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 18 | 6.1 | 6 | 7 | 0.539 | 0.5 | 0.55 | 0.214 | 0.177 | 0.2103
[23 empty] | Indices | 6 | 7 |  | 4 | 5 |  |  |  |  | Indices | 6 | 7 |  | 4 | 5 | [16 empty] | 0.214 | 0.177 |  | 0.214 | 0.177
[22 empty] | 18 | 6.1 | 6 | 7 | 0.539 | 0.5 | 0.55 | 1.071 | 1.135 | 1.0774 | 18 | 6.1 | 6 | 7 | 0.539 | 0.5 | 0.55 | 1.144 | 1.204 | 1.15 | [12 empty] | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.071 | 1.135 |  | 1.071 | 1.135 | [5 empty] | 1.144 | 1.204 |  | 1.144 | 1.204 | [14 empty] | 19 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2316
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 19 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0442 | 19 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1164 | [12 empty] | Indices | 6 | 7 |  | 4 | 5
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144 | [14 empty] | 20 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 0.258 | 0.214 | 0.2316
[23 empty] | Indices | 5 | 6 |  | 4 | 5 |  |  |  |  | Indices | 5 | 6 |  | 4 | 5 | [16 empty] | 0.258 | 0.214 |  | 0.258 | 0.214
[22 empty] | 20 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.004 | 1.071 | 1.0442 | 20 | 5.6 | 5 | 6 | 0.539 | 0.5 | 0.55 | 1.075 | 1.144 | 1.1164
[24 empty] | 1.004 | 1.071 |  | 1.004 | 1.071 | [5 empty] | 1.075 | 1.144 |  | 1.075 | 1.144
Note on B17: (mm Al) measured behind compression plate
Note on B21: fpro: Difference between real and measured compressed breast thickness
Note on D25: Measured or showed on display
Note on E7: Fill only cells in blue color
Note on G21: Real thickness-thickness showed on display
Note on G23: fisica1: fr is HVL dependent
Note on G25: W-Rh Rh-Al

## PMMA dosimetry_1
Dose to typical breasts simulated with PMMA
[11 empty] | Factor s |  |  | ánodo/filtro |  |  | Espesor de filtro |  | n | a | b | A | c
[11 empty] | Mo-Mo | 1 |  |  | Mo/30 mm Mo |  | 36,1 mm |  | 3.06 | -0.000326 | 0.0273 | 0.0010662 | -0.13672 | Mo-Mo
Date | [10 empty] | Mo-Rh | 1.017 |  |  | Mo/25 mm Rh |  | 29,9 mm |  | 3.24 | -0.000624 | 0.0445 | 0.00047602 | -0.32773 | Mo-Rh
[11 empty] | Rh-Rh | 1.061 |  |  | Rh/25 mm Rh |  | 29,9 mm |  | 3.03 | -0.000514 | 0.0425 | 0.00093951 | -0.33839 | Rh-Rh
Facility:
Room: | [10 empty] | Factor c PMMA | [10 empty] | Valor de g PMMA
Unit: | [10 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
[9 empty] | kVp set | kVpmeas | Espesor | HVL (mm Al) | [9 empty] | Valor de g PMMA 1: Espesor; Valor de g PMMA 2: HVL (mm Al)
[9 empty] | 23 | 23 | mama | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  |  |  | PMMA | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
[9 empty] | 24 | 24 | 2 | 0.889 | 0.895 | 0.903 | 0.908 | 0.912 | 0.917 | 0.921 |  |  |  | 2 | 0.329 | 0.378 | 0.421 | 0.46 | 0.496 | 0.529 | 0.559 | 0.585
[9 empty] | 25 | 25 | 3 | 0.94 | 0.943 | 0.945 | 0.946 | 0.949 | 0.952 | 0.953 |  |  |  | 3 | 0.222 | 0.261 | 0.294 | 0.326 | 0.357 | 0.388 | 0.419 | 0.448
|  |  |  | Mo-Mo | Mo-Rh | Rh-Rh |  |  | 26 | 26 | 4 | 1.043 | 1.041 | 1.04 | 1.039 | 1.037 | 1.035 | 1.034 |  |  |  | 4 | 0.155 | 0.183 | 0.208 | 0.232 | 0.258 | 0.285 | 0.311 | 0.339
kV ( nominal) |  |  |  | 28 | 28 | 28 |  |  | 27 | 27 | 4.5 | 1.109 | 1.105 | 1.102 | 1.099 | 1.096 | 1.091 | 1.088 |  |  |  | 4.5 | 0.13 | 0.155 | 0.177 | 0.198 | 0.22 | 0.245 | 0.272 | 0.295
kVmeas at kVnom |  |  |  | 28.1 | 28.1 | 28.1 |  |  | 28 | 28 | 5 | 1.164 | 1.16 | 1.151 | 1.15 | 1.144 | 1.139 | 1.134 |  |  |  | 5 | 0.112 | 0.135 | 0.154 | 0.172 | 0.192 | 0.214 | 0.236 | 0.261
Output at kVnom |  |  |  | 28.9 | 23.52 | 23.04 |  |  | 29 | 29 | 6 | 1.254 | 1.245 | 1.235 | 1.231 | 1.225 | 1.217 | 1.207 |  |  |  | 6 | 0.088 | 0.106 | 0.121 | 0.136 | 0.152 | 0.166 | 0.189 | 0.21
HVL at kVnom |  |  |  | 0.373 | 0.43 | 0.45 |  |  | 30 | 30 | 7 | 1.299 | 1.292 | 1.282 | 1.275 | 1.27 | 1.26 | 1.249 |  |  |  | 7 |  | 0.086 | 0.098 | 0.111 | 0.123 | 0.136 | 0.154 | 0.172
[9 empty] | 31 | 31
Focus-platform distance (cm): |  |  |  | 63.7 |  |  |  |  | 32 | 32
[9 empty] | 33 | 33
MGD = | K·c.g·s | [7 empty] | 34 | 34
[12 empty] | Interpolación c | [9 empty] | Valor de g PMMA 1: Interpolación g
Thickness (cm) | kVps | Target/Filter | mAs |  |  |  | ESAK (mGy) | MGD (mGy) |  |  |  | Indices |  | HVL1 | HVL2 | c1 | c2 |  |  |  |  | Indices |  | HVL1 | HVL2 | g1 | g2 |  |  | HVL | KVpm | DFPiel (cm) | Rend. 1m (mGy/mAs) | Factor c | Factor g | Factor s
2 | 25 | Mo-Mo | 17 |  |  |  | 0.90247 | 0.33415 |  |  |  | 2 | 3 | 0.3 | 0.35 | 0.889 | 0.895 |  |  |  |  | 3 | 4 | 0.3 | 0.35 | 0.378 | 0.421 |  |  | 0.34203 | 25 | 61.7 | 20.2094 | 0.89404 | 0.41415 | 1
3 | 26 | Mo-Mo | 30 |  |  |  | 1.8553 | 0.51746 |  |  |  | 3 | 4 | 0.35 | 0.4 | 0.943 | 0.945 |  |  |  |  | 4 | 5 | 0.35 | 0.4 | 0.294 | 0.326 |  |  | 0.35271 | 26 | 60.7 | 22.7864 | 0.94311 | 0.29573 | 1
4 | 28 | Mo-Mo | 42 |  |  |  | 3.3687 | 0.76629 |  |  |  | 3 | 4 | 0.35 | 0.4 | 1.041 | 1.04 |  |  |  |  | 4 | 5 | 0.35 | 0.4 | 0.208 | 0.232 |  |  | 0.3721 | 28 | 59.7 | 28.5864 | 1.0406 | 0.21861 | 1
4.5 | 27 | Mo-Rh | 58 |  |  |  | 3.42 | 0.78994 |  |  |  | 4 | 5 | 0.4 | 0.45 | 1.102 | 1.099 |  |  |  |  | 5 | 6 | 0.4 | 0.45 | 0.198 | 0.22 |  |  | 0.41887 | 27 | 59.2 | 20.6656 | 1.1009 | 0.2063 | 1.017
5 | 27 | Mo-Rh | 84 |  |  |  | 5.0379 | 1.0585 |  |  |  | 4 | 5 | 0.4 | 0.45 | 1.151 | 1.15 |  |  |  |  | 5 | 6 | 0.4 | 0.45 | 0.172 | 0.192 |  |  | 0.41887 | 27 | 58.7 | 20.6656 | 1.1506 | 0.17955 | 1.017
6 | 31 | Mo-Rh | 48 |  |  |  | 4.6615 | 0.89031 |  |  |  | 5 | 6 | 0.45 | 0.5 | 1.231 | 1.225 |  |  |  |  | 6 | 7 | 0.45 | 0.5 | 0.152 | 0.166 |  |  | 0.4521 | 31 | 57.7 | 32.3326 | 1.2307 | 0.15259 | 1.017
7 | 32 | Mo-Rh | 84 |  |  |  | 9.3633 | 1.5155 |  |  |  | 5 | 6 | 0.45 | 0.5 | 1.275 | 1.27 |  |  |  |  | 6 | 7 | 0.45 | 0.5 | 0.123 | 0.136 |  |  | 0.45729 | 32 | 56.7 | 35.8356 | 1.2743 | 0.1249 | 1.017
Note on B17: (mm Al) measured behind compression plate
Note on D25: Mo-Mo Mo-Rh Rh-Rh
Note on E4: Fill only cells in blue color

## PMMA dosimetry_2
Dose to typical breasts simulated with PMMA
[10 empty] | Factor s |  |  | ánodo/filtro |  |  | Espesor de filtro |  | n | a | b | A | c
[10 empty] | W-Rh | 1.042 |  | W/50 mm Rh |  | 58,9 mm |  |  | 1.96 | -0.00054 | 0.0403 | 0.041825 | -0.16604 | W-Rh
Date | [9 empty] | Rh-Al | 1.044 |  | Rh/1,0 mm Al |  | 1,2 mm |  |  | 4.39 | -0.00113 | 0.0909 | 1.0271e-05 | -1.102 | Rh-Al
[13 empty] | Mo/1,0 mm Al |  | 1,2 mm |  |  | 4.23 | -0.000775 | 0.0593 | #DIV/0! | -1.0544 | Mo-Al
Facility:
Room: | [9 empty] | Factor c PMMA | [10 empty] | Valor de g PMMA
Unit: | [9 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
[8 empty] | kVp set | kVpmeas | Espesor | HVL (mm Al) | [9 empty] | Valor de g PMMA 1: Espesor; Valor de g PMMA 2: HVL (mm Al)
[8 empty] | 23 | 23 | mama | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  |  |  | PMMA | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
[8 empty] | 24 | 24 | 2 | 0.889 | 0.895 | 0.903 | 0.908 | 0.912 | 0.917 | 0.921 |  |  |  | 2 | 0.329 | 0.378 | 0.421 | 0.46 | 0.496 | 0.529 | 0.559 | 0.585
[8 empty] | 25 | 25 | 3 | 0.94 | 0.943 | 0.945 | 0.946 | 0.949 | 0.952 | 0.953 |  |  |  | 3 | 0.222 | 0.261 | 0.294 | 0.326 | 0.357 | 0.388 | 0.419 | 0.448
|  |  |  | W-Rh | Rh-Al |  |  | 26 | 26 | 4 | 1.043 | 1.041 | 1.04 | 1.039 | 1.037 | 1.035 | 1.034 |  |  |  | 4 | 0.155 | 0.183 | 0.208 | 0.232 | 0.258 | 0.285 | 0.311 | 0.339
kV ( nominal) |  |  |  | 28 | 28 |  |  | 27 | 27 | 4.5 | 1.109 | 1.105 | 1.102 | 1.099 | 1.096 | 1.091 | 1.088 |  |  |  | 4.5 | 0.13 | 0.155 | 0.177 | 0.198 | 0.22 | 0.245 | 0.272 | 0.295
kVmeas at kVnom |  |  |  | 28.1 | 28.1 |  |  | 28 | 28 | 5 | 1.164 | 1.16 | 1.151 | 1.15 | 1.144 | 1.139 | 1.134 |  |  |  | 5 | 0.112 | 0.135 | 0.154 | 0.172 | 0.192 | 0.214 | 0.236 | 0.261
Output at kVnom |  |  |  | 28.9 | 23.52 |  |  | 29 | 29 | 6 | 1.254 | 1.245 | 1.235 | 1.231 | 1.225 | 1.217 | 1.207 |  |  |  | 6 | 0.088 | 0.106 | 0.121 | 0.136 | 0.152 | 0.166 | 0.189 | 0.21
HVL at kVnom |  |  |  | 0.54 | 0.56 |  |  | 30 | 30 | 7 | 1.299 | 1.292 | 1.282 | 1.275 | 1.27 | 1.26 | 1.249 |  |  |  | 7 |  | 0.086 | 0.098 | 0.111 | 0.123 | 0.136 | 0.154 | 0.172
[8 empty] | 31 | 31
Focus-platform distance (cm): |  |  |  | 63.7 |  |  |  | 32 | 32
[8 empty] | 33 | 33
MGD = | K·c.g·s | [6 empty] | 34 | 34
[11 empty] | Interpolación c | [9 empty] | Valor de g PMMA 1: Interpolación g
Thickness (cm) | kVps | Target/Filter | mAs |  |  | ESAK (mGy) | MGD (mGy) |  |  |  | Indices |  | HVL1 | HVL2 | c1 | c2 |  |  |  |  | Indices |  | HVL1 | HVL2 | g1 | g2 |  |  | HVL | KVpm | DFPiel (cm) | Rend. 1m (mGy/mAs) | Factor c | Factor g | Factor s
2 | 25 | W-Rh | 17 |  |  | 1.0263 | 0.51847 |  |  |  | 6 | 7 | 0.5 | 0.55 | 0.912 | 0.917 |  |  |  |  | 7 | 8 | 0.5 | 0.55 | 0.529 | 0.559 |  |  | 0.50396 | 25 | 61.7 | 22.9824 | 0.9124 | 0.53138 | 1.042
3 | 26 | W-Rh | 30 |  |  | 2.0208 | 0.79689 |  |  |  | 6 | 7 | 0.5 | 0.55 | 0.949 | 0.952 |  |  |  |  | 7 | 8 | 0.5 | 0.55 | 0.388 | 0.419 |  |  | 0.51672 | 26 | 60.7 | 24.8188 | 0.95 | 0.39837 | 1.042
4 | 28 | Rh-Al | 42 |  |  | 2.7286 | 0.92878 |  |  |  | 7 | 8 | 0.55 | 0.6 | 1.035 | 1.034 |  |  |  |  | 8 | 9 | 0.55 | 0.6 | 0.311 | 0.339 |  |  | 0.55725 | 28 | 59.7 | 23.1548 | 1.0349 | 0.31506 | 1.044
4.5 | 27 | Rh-Al | 58 |  |  | 3.2665 | 0.97072 |  |  |  | 6 | 7 | 0.5 | 0.55 | 1.096 | 1.091 |  |  |  |  | 7 | 8 | 0.5 | 0.55 | 0.245 | 0.272 |  |  | 0.5285 | 27 | 59.2 | 19.738 | 1.0932 | 0.26039 | 1.044
5 | 27 | Rh-Al | 84 |  |  | 4.8118 | 1.2987 |  |  |  | 6 | 7 | 0.5 | 0.55 | 1.144 | 1.139 |  |  |  |  | 7 | 8 | 0.5 | 0.55 | 0.214 | 0.236 |  |  | 0.5285 | 27 | 58.7 | 19.738 | 1.1412 | 0.22654 | 1.044
6 | 26 | Rh-Al | 48 |  |  | 2.4112 | 0.50986 |  |  |  | 5 | 6 | 0.45 | 0.5 | 1.231 | 1.225 |  |  |  |  | 6 | 7 | 0.45 | 0.5 | 0.152 | 0.166 |  |  | 0.49749 | 26 | 57.7 | 16.7244 | 1.2253 | 0.1653 | 1.044
7 | 28 | Rh-Al | 84 |  |  | 6.05 | 1.2448 |  |  |  | 7 | 8 | 0.55 | 0.6 | 1.26 | 1.249 |  |  |  |  | 8 | 9 | 0.55 | 0.6 | 0.154 | 0.172 |  |  | 0.55725 | 28 | 56.7 | 23.1548 | 1.2584 | 0.15661 | 1.044
Note on B17: (mm Al) measured behind compression plate
Note on D25: Mo-Mo Mo-Rh Rh-Rh
Note on E4: Fill only cells in blue color