## Method
CT Scan Protocol for Paediatric Patients
Based on " A method for the systematic selection of technique factors in paediatric CT (BJR 76 (2003), 51-56.
Authors: | Maria Angeles Rivas Ballarín and Pedro Ruiz Manzano
Hospital Clínico Universitario "Lozano Blesa"
Avda San Juan Bosco 15. 50009-Zaragoza .Spain.
mrivasb@salud.aragon.es
pruizm@salud.aragon.es
The approach is based on the following assumptions:
1. The technique factors for the equivalent adult examination are known.
2. The image signal-to-noise ratio in the equivalent adult examination is suitable for diagnosis in the paediatric case.
3. The attenuation and relative proportions of body tissues are similar in adults and paediatrics.
4. The tube voltage is not altered between the adult and paediatric versions of the examination.
5. The reconstruction filter used in the paediatric examination is not radically different from that used in the adult one.
Patient cross-sectional size is approximated from the equivalent diameter, d. This is obtained measuring the perimeter of the section to be scaned and dividing by p.
The following equation is used:
Where
SP, SA are the slice width for children and adult.
dA , dP are the standard adult and paediatric equivalent diameter respectively.
µeff is the effective linear attenuation coefficient for the body segment being scanned at the tube potential being used.

## Head
Paediatric Protocols: Head CT
Fill only cells in green colour.
Child's head perimeter |  | 36.5 |  |  | Datos equipo |  |  | meff cabeza
[5 empty] | Filtración (mm Al) = | 12 | 1 | kVp/F | 9 | 12
Adul Protocol |  |  | Paediatric Protocol |  |  |  | 2 | 120 | 0.194 | 0.192
kVp = | 140 |  | kVp = | 140 | Diámetro adulto dA = | 17.6 | 3 | 140 | 0.186 | 0.185
mAsA = | 309 |  | mAsP = | 102.1788 | Diámetro niñodP = | 11.6183
Slice width SA (mm) = | 8 |  | Slice width SP (mm) = | 8 |  |  |  | Indica fila | meff
[8 empty] | 3 | 0.185
WATCH OUT! If mAsP vaule is below the minimum mAs available, slice width should be reduced.
|  | New slice width S'P (mm) = |  | 5
|  | New mAsP value = |  | 163.486

## Chest
Paediatric Protocols: Chest CT
Fill only cells in green colour.
Child's chest perimeter |  | 73 |  |  | Datos equipo |  |  | meff tórax
[5 empty] | Filtración (mm Al) = | 12 | 1 | kVp/F | 9 | 12
Adul Protocol |  |  | Paediatric Protocol |  |  |  | 2 | 120 | 0.146 | 0.145
kVp = | 120 |  | kVp = | 120 | Diámetro adulto dA = | 27.6 | 3 | 140 | 0.14 | 0.138
mAsA = | 180 |  | mAsP = | 95.6083 | Diámetro niñodP = | 23.2366
Slice width SA (mm) = | 8 |  | Slice width SP (mm) = | 8 |  |  |  | Indica fila | meff
[8 empty] | 2 | 0.145
WATCH OUT! If mAsP vaule is below the minimum mAs available, slice width should be reduced.
|  | New slice width S'P (mm) = |  | 5
|  | New mAsP value = |  | 152.9732

## Abdomen
Paediatric Protocols: Abdomen
Fill only cells in green colour.
Child's abdomen perimeter |  | 63 |  |  | Datos equipo |  |  | meff abdomen
[5 empty] | Filtración (mm Al) = | 12 | 1 | kVp/F | 9 | 12
Adul Protocol |  |  | Paediatric Protocol |  |  |  | 2 | 120 | 0.184 | 0.184
kVp = | 120 |  | kVp = | 120 | Diámetro adulto dA = | 23.4 | 3 | 140 | 0.177 | 0.176
mAsA = | 211 |  | mAsP = | 113.9885 | Diámetro niñodP = | 20.0535
Slice width SA (mm) = | 8 |  | Slice width SP (mm) = | 8 |  |  |  | Indica fila | meff
[8 empty] | 2 | 0.184
WATCH OUT! If mAsP vaule is below the minimum mAs available, slice width should be reduced.
|  | New slice width S'P (mm) = |  | 3
|  | New mAsP value = |  | 303.9694